## Slide 1: Heating and Stirring Module Design Review #4

## Slide 2: Power Constraints
Underestimated power requirement of hotplates.
  10-25W required, 5W initial estimate.
  Further testing required.
  Total power consumption, 450-700W.
Change from 5V to 24V architecture.
  Lower transmission losses.
  Lower current.
  Requires redesign of current power distribution systems.

## Slide 3: Power contd.
Larger power supply required, increased cost.
  LRS-600N2, 600W @ 24V, $64. 12% larger footprint.
  Cannot accommodate a 6-module control hub, 4 at most.
H&S 24-5V Buck required, increased complexity.

## Slide 4: Redesigned Hotplate
Tuned line thickness, routing pattern for:
  Higher resistance.
  Output power, 25W goal.
  Heat dissipation.
  Thermistor accuracy.

## Slide 5: Adjusted Housing Dimensions
Altered motor spacing to accommodate for a Nucleo-32 development board.
Reduced complexity.
All housing components must be updated.

## Slide 6: Heating and Stirring PCB
75% complete, missing pinouts, several connectors.
4-layers, assembled by hand.
SMD components unavoidable.
Redesigned power distribution.
  24V, 1A hotplate. Main bus.
  5V, ~3A total motors. Buck converter.
  3.3V, ~400mA logic. Internal LDO VR.

## Slide 7: Purchasing PCB’s
Low materials cost
Customs fee & tariffs ~$40
Ordering in bulk
Extensive testing prior to order

## Slide 8: Next Steps
Finish H&S and Control Hub PCBs.
Remodel H&S lid and vial guide.
Remodel Control Hub.
Test stir bar interaction.